Micro-CT Ring Removal
User Manual
Current GUI version

Purpose
This manual describes how to use the graphical interface for ring artefact reduction and, when needed, reconstruction through CERA. It focuses on the controls visible in the GUI and on the most common processing scenarios.

# 1. What the software does
Micro-CT Ring Removal is a graphical tool for processing micro-CT projection data. It can be used in three ways: ring removal only, reconstruction only, or ring removal followed by reconstruction.
One or more projection folders can be queued in a single session. Each folder is processed as a separate job. All ring-removal settings active in the GUI at the moment the run starts are applied to every queued job in that run.
Reconstruction is performed through an external CERA Python environment. For routine work, the program automatically uses the reconstruction config file found in each selected input folder. A manual config override is available for special cases.
# 2. Main window overview
The main window is divided into the sections listed below.
Section	Purpose
Input projection folders	Add and manage folders that contain the projection images to be processed.
Processing pipeline	Choose whether the run performs ring removal only, reconstruction only, or both steps in sequence.
Output	Choose where corrected projections will be written when ring removal is part of the pipeline.
Ring-removal algorithm	Select the correction mode, the correction method, and the method-specific parameters.
Reconstruction (CERA)	Select the CERA Python executable, define a reconstruction name, and optionally override automatic config selection.
Run	Start the queued jobs and monitor the progress bar.
Log	Review status messages, saved output paths, and error details.

# 3. Input projection folders
This section defines which datasets will be processed.
Control	Description
Add folder(s)	Adds one projection folder to the processing list.
Remove selected	Removes the highlighted folder from the list.
Clear	Removes all folders from the list.
Glob pattern	Defines which files inside each folder are treated as input images.
Recursive	Searches subfolders when enabled.

The default glob pattern is tomo_*.tif. This matches file names such as tomo_0001.tif, tomo_0002.tif, and similar TIFF sequences.
## How to use this section
1. Click Add folder(s).
2. Select a folder that contains the projection TIFF files you want to process.
3. Repeat the step if you want to queue several folders in one run.
4. Change the Glob pattern if your files use a different naming scheme.
5. Enable Recursive only when the projection images are stored in nested subfolders.
# 4. Processing pipeline
The Processing pipeline section decides which operations will be carried out for each selected folder.
Option	What the program does
Ring removal only	Creates corrected projections and stops there.
Reconstruction only	Runs CERA directly from the selected input folder without creating a corrected-projection folder.
Ring removal + reconstruction	First creates corrected projections, then runs CERA using the corrected projections as input.

When Reconstruction only is selected, the normal Output section is not used for reconstruction output. CERA writes the reconstruction results in the same folder that contains the original projections and the corresponding config file.
When Ring removal + reconstruction is selected, CERA writes the reconstruction results in the folder that contains the corrected projections.
# 5. Output section
The Output section controls where corrected projections are written when ring removal is part of the pipeline.
Option	Description
Save mode	Chooses the output location strategy.
Created folder name	Sets the name of the output folder created for corrected projections.
Custom output / Browse...	Lets you choose an exact destination when Save mode is set to Custom folder.
Overwrite output files	Allows existing corrected output files to be replaced.
Keep temporary sinogram files	Preserves temporary sinogram files after processing.

## Available save modes
Save mode	Meaning
Inside input folder	Creates the corrected result inside the selected input folder.
One level above input folder	Creates the corrected result in the parent folder above the selected input folder.
One level below / child folder	Creates a child folder below the selected input folder.
Custom folder	Saves corrected output to a manually selected destination.

The default created folder name is ring_corrected.
The Custom output field and Browse... button are active only when Save mode is set to Custom folder.
For routine work, One level below / child folder with the default folder name ring_corrected is a practical choice. It keeps corrected data close to the raw projections while keeping the output separate from the original files.
# 6. Ring-removal algorithm section
This section contains the controls for correction mode, correction method, parameter values, and worker count.
## Mode
The available Mode options are auto, intensity, and log.
For most users, auto is the best starting point. Use intensity or log only when you already know the data domain and want direct control over the processing path.
## Correction method
The available Correction methods are auto, algotom, repair, filtering, sorting, wavelet_fft, dead, and large. The default method is algotom. When the selected method changes, the Parameters panel updates so that only the relevant settings are shown.
## Method-specific parameters
Method	Parameters shown in the GUI
auto	Displays a note explaining that algotom is used for intensity sinograms and repair for log sinograms.
algotom	SNR, la_size, sm_size, dim
repair	repair_thresh, repair_max_cols
filtering	filt_sigma, filt_size, filt_dim, filt_sort
sorting	sort_size, sort_dim
wavelet_fft	wfft_level, wfft_size, wfft_wavelet_name, wfft_window_name, wfft_sort
dead	dead_snr, dead_size, dead_residual
large	large_snr, large_size, large_drop_ratio, large_norm

## Workers
Workers controls how many worker processes are used during correction. The GUI allows values from 0 to 128, with a default of 12. A value of 0 means automatic selection.
## Practical guidance
- Mode: auto
- Correction: algotom
- Workers: 12 or fewer, depending on available system resources
- Keep the method-specific parameters at their defaults for the first test run
If the first result is not satisfactory, test one correction method at a time and change only one or two parameters between runs. This makes it easier to identify which change improved or degraded the result.
# 7. Reconstruction (CERA) section
This section is used whenever the selected pipeline includes reconstruction.
Control	Description
CERA Python executable	Path to the Python executable in the environment where the CERA wheel is installed.
Use custom CERA config template	When enabled, lets you manually select a config template instead of using the default config discovered in each input folder.
Custom CERA config template	Manual override used only when the custom option is enabled.
Reconstruction name	Name used when rendering the reconstruction config; if left empty, the input folder name is used.

## Default config selection
- Reconstruction only: the program looks for the reconstruction config in the selected input folder and uses it automatically.
- Ring removal + reconstruction: the program takes the config from the original input folder, renders a working config in the corrected-projection folder, and runs CERA there.
For routine work, the custom config override should be left disabled. It is intended for cases where more than one config is present or where a non-standard template must be used.
## Important note about the Python executable
The Python executable selected in this section must belong to the CERA environment, not to the main GUI environment. If the wrong interpreter is selected, reconstruction will fail even if ring removal works correctly.
# 8. Run section
The Run section contains the Run selected jobs button and a progress bar.
When Run selected jobs is clicked, the program checks that at least one input folder has been added, validates the settings needed for the selected pipeline, builds the job list from the queued folders, and starts one worker per folder.
The progress bar tracks the major processing stages. The exact number of stages depends on the selected pipeline: ring removal only, reconstruction only, or the combined workflow.
# 9. Log section
The Log section is a read-only text panel that displays processing messages.
- Queued folders
- Progress updates
- Saved output locations
- Saved text log and JSON summary paths
- Reconstruction messages from CERA
- Errors and tracebacks
When a job succeeds, the log reports the output directory. If reconstruction is part of the pipeline, the log also reports the reconstruction output path and the rendered config path.
Each run also creates a text log and a JSON summary in the corresponding output folder. These files are useful for checking the exact settings used during a run.
# 10. Typical workflows
## Scenario 1: Quick standard correction for one dataset
1. Add the projection folder.
2. Leave the Glob pattern as tomo_*.tif if the files use that naming style.
3. In Output, choose One level below / child folder.
4. Keep Created folder name as ring_corrected, or change it if needed.
5. In Ring-removal algorithm, set Mode = auto and Correction = algotom.
6. Set Workers to 12 or lower if your system requires it.
7. Click Run selected jobs.
8. Monitor the progress bar and log messages.
9. Inspect the corrected output folder when processing is complete.
## Scenario 2: Reconstruction only from original projections
1. Add the projection folder.
2. Set Processing pipeline to Reconstruction only.
3. In Reconstruction (CERA), select the Python executable from the CERA environment.
4. Leave Use custom CERA config template disabled for normal operation.
5. If needed, enter a reconstruction name. If left empty, the input folder name is used.
6. Click Run selected jobs.
7. Review the log. The reconstruction output is written in the same folder as the original projections and config file.
## Scenario 3: Ring removal followed by reconstruction
1. Add the projection folder.
2. Set Processing pipeline to Ring removal + reconstruction.
3. Configure the Output section for the corrected projections.
4. Set the ring-removal parameters.
5. In Reconstruction (CERA), select the Python executable from the CERA environment.
6. Leave Use custom CERA config template disabled unless a manual override is required.
7. Click Run selected jobs.
8. The program creates corrected projections first, then runs reconstruction from the corrected-projection folder.
## Scenario 4: Batch processing of several folders
1. Add all required folders to the input list.
2. Choose the processing pipeline.
3. Choose the output settings if ring removal is part of the pipeline.
4. Select the ring-removal method and parameter set if needed.
5. Click Run selected jobs.
Each folder becomes a separate job. If reconstruction is included, each folder uses its own config file from that folder unless the custom override is enabled.
## Scenario 5: Files are stored in subfolders
1. Add the main folder.
2. Enable Recursive.
3. Confirm that the Glob pattern matches the projection files.
4. Run the job normally.
## Scenario 6: Compare different correction methods
1. Process the same dataset multiple times.
2. Use a different output folder name for each run, such as ring_corrected_algotom, ring_corrected_wavelet, or ring_corrected_filtering.
3. Change only one major setting between runs.
4. Compare the outputs visually.
# 11. Suggestions for choosing methods
A practical starting order for most cases is:
- algotom
- wavelet_fft
- filtering
- sorting
- dead or large for more specific stripe behaviour
- repair when working in log-type conditions or when testing the dedicated repair path
- auto when you want the program to choose between algotom and repair based on the detected domain
This order is a usage recommendation based on the options available in the GUI. It is not a fixed rule enforced by the program.
# 12. Troubleshooting
Issue	Action
No input added	Add at least one projection folder before clicking Run selected jobs.
Nothing found in folder	Check the Glob pattern first. The default is tomo_*.tif.
Output already exists	Enable Overwrite output files only when you intend to replace older results.
Job fails during reconstruction	Check that the CERA Python executable points to the environment where the CERA wheel is installed.
No config found for reconstruction	Make sure the input folder contains the expected .config or .cfg file, or enable the custom config override.
Too many workers	Reduce the Workers value if the computer becomes slow or unstable. The default is 12; 0 means automatic.

# 13. Everyday quick-start
- Add one or more projection folders.
- Leave the Glob pattern as tomo_*.tif unless your files use a different name.
- Choose the processing pipeline.
- For ring removal, start with Mode = auto and Correction = algotom.
- For reconstruction, select the Python executable from the CERA environment.
- Leave the custom config override disabled unless it is really needed.
- Run the job.
- Read the log and review the resulting files.

Default GUI values referenced in this manual: Correction = algotom; Workers = 12.