Micro-CT Ring Removal
User Manual
Current GUI version

Purpose This manual explains how to use the graphical interface to reduce ring artefacts in micro-CT projection data, from selecting input folders to reviewing job progress and output messages.

# 1. What the software does
Micro-CT Ring Removal is a graphical tool for reducing ring artefacts in micro-CT projection data based on algotom algorithms (https://algotom.readthedocs.io/en/latest/). The program lets you select one or more folders containing projection images, choose an output location, define correction settings, and run the selected jobs from a single interface.
Each selected input folder is processed as a separate job. All jobs in the same run use the algorithm and output settings that are active in the GUI when processing starts.
# 2. Main window overview
The main window is divided into five sections:
Section	Purpose
Input projection folders	Add and manage the folders that contain projection images to be corrected.
Output	Choose where corrected data will be saved and how the output folder is created.
Algorithm	Select the correction mode, the correction method, and any method-specific parameters.
Run	Start the selected jobs and monitor the progress bar.
Log	Review status messages, success notices, and error details.

# 3. Input projection folders
The Input projection folders section defines which datasets will be processed.
Control	Description
Add folder(s)	Adds one projection folder to the processing list.
Remove selected	Removes the highlighted folder from the list.
Clear	Removes all folders from the list.
Glob pattern	Defines which files inside the folder are treated as input images.
Recursive	Searches subfolders when enabled.

The default glob pattern is tomo_*.tif. This matches file names such as tomo_0001.tif, tomo_0002.tif, and similar TIFF sequences.
## How to use this section
1. Click Add folder(s).
2. Select a folder that contains the projection TIFF files you want to process.
3. Repeat the step if you want to queue several folders in one run.
4. Change the Glob pattern if your files use a different naming scheme.
5. Enable Recursive only when the projection images are stored in nested subfolders.
# 4. Output section
The Output section controls where corrected results are written.
Option	Description
Save mode	Chooses the output location strategy.
Created folder name	Sets the name of the output folder created for corrected data.
Custom output / Browse...	Lets you choose an exact destination when Save mode is set to Custom folder.
Overwrite output files	Allows existing output files to be replaced.
Keep temporary sinogram files	Preserves temporary sinogram files after processing.

## Available save modes
Save mode	Meaning
Inside input folder	Creates the corrected result inside the selected input folder.
One level above input folder	Creates the corrected result in the parent folder above the selected input folder.
One level below / child folder	Creates a child folder below the selected input folder.
Custom folder	Saves output to a manually selected destination.

The default created folder name is ring_corrected.
The Custom output field and Browse... button are active only when Save mode is set to Custom folder.
## Recommended practice
For routine work, One level below / child folder with the default folder name ring_corrected is a practical choice. It keeps corrected data close to the raw projections while keeping the output separate from the original files. Leave Overwrite output files disabled unless you intentionally want to replace an earlier result.
# 5. Algorithm section
The Algorithm section contains the controls for correction mode, correction method, parameter values, and worker count.
## Mode
The available Mode options are auto, intensity, and log.
For most users, auto is the best starting point. Use intensity or log only when you already know the data domain and want direct control over the processing path.
## Correction method
The available Correction methods are auto, algotom, repair, filtering, sorting, wavelet_fft, dead, and large. The default method is algotom.
When you change the selected method, the Parameters panel updates so that only the relevant settings are shown.
## Method-specific parameters
Method	Parameters shown in the GUI
auto	Displays a note explaining that algotom is used for intensity sinograms and repair for log sinograms.
algotom	SNR, la_size, sm_size, dim
repair	repair_thresh, repair_max_cols
filtering	filt_sigma, filt_size, filt_dim, filt_sort
sorting	sort_size, sort_dim
wavelet_fft	wfft_level, wfft_size, wfft_wavelet_name, wfft_window_name, wfft_sort
dead	dead_snr, dead_size, dead_residual
large	large_snr, large_size, large_drop_ratio, large_norm

## Workers
Workers controls how many worker processes are used during correction. The GUI allows values from 0 to 128, with a default of 12. A value of 0 means automatic selection.
## Practical guidance
- Mode: auto
- Correction: algotom
- Workers: 12 or fewer, depending on available system resources
- Keep the method-specific parameters at their defaults for the first test run
If the first result is not satisfactory, test one correction method at a time and change only one or two parameters between runs. This makes it easier to identify which change improved or degraded the result.
# 6. Run section
The Run section contains the Run selected jobs button and a progress bar.
When Run selected jobs is clicked, the program verifies that at least one input folder has been added, builds the job list from the queued folders, applies one shared parameter set from the current GUI settings, queues one worker per folder, and disables the run button until all queued jobs have finished.
The progress bar covers three major workflow stages, which is why it is configured from 0 to 3.
# 7. Log section
The Log section is a read-only text panel that displays processing messages.
- Queued folders
- Progress updates
- Successful completion messages
- Errors and tracebacks
When a job succeeds, the output directory is reported in the log. When a job fails, the log shows ERROR followed by the traceback text returned by the worker. After all queued jobs are complete, the program displays a final message confirming that processing has finished.
# 8. Typical workflows
## Scenario 1: Quick standard correction for one dataset
1. Add the projection folder.
2. Leave the Glob pattern as tomo_*.tif if the files use that naming style.
3. In Output, choose One level below / child folder.
4. Keep Created folder name as ring_corrected, or change it if needed.
5. In Algorithm, set Mode = auto and Correction = algotom.
6. Set Workers to 12 or lower if your system requires it.
7. Click Run selected jobs.
8. Monitor the progress bar and log messages.
9. Inspect the corrected output folder when processing is complete.
## Scenario 2: Batch processing of several folders
- Add all required folders to the input list.
- Choose the output settings once.
- Select the correction method and parameter set.
- Click Run selected jobs.
Each folder becomes a separate job, but every job uses the same algorithm and output settings that were active when the run started.
## Scenario 3: Files are stored in subfolders
- Add the main folder.
- Enable Recursive.
- Confirm that the Glob pattern matches the projection files.
- Run the job normally.
## Scenario 4: Compare different correction methods
- Process the same dataset multiple times.
- Use a different output folder name for each run, such as ring_corrected_algotom, ring_corrected_wavelet, or ring_corrected_filtering.
- Change only one major setting between runs.
- Compare the outputs visually.
## Scenario 5: Keep intermediate files for inspection
- Enable Keep temporary sinogram files.
- Run the job.
- Inspect the preserved temporary files after processing.
# 9. Suggestions for choosing methods
A practical starting order for most cases is:
algotom
wavelet_fft
filtering
sorting
dead or large for more specific stripe behaviour
repair when working in log-type conditions or when testing the dedicated repair path
auto when you want the program to choose between algotom and repair based on the detected domain
This order is a usage recommendation based on the options available in the GUI and the method names presented in the interface. It is not a fixed rule enforced by the program.
# 10. Troubleshooting
Issue	Action
No input added	Add at least one projection folder before clicking Run selected jobs.
Nothing found in folder	Check the Glob pattern first. The default is tomo_*.tif.
Output already exists	Enable Overwrite output files only when you intend to replace older results.
Job fails	Read the Log panel. Error tracebacks are shown there when an exception occurs.
Too many workers	Reduce the Workers value if the computer becomes slow or unstable. The default is 12; 0 means automatic.

# 11. Everyday quick-start
For routine use, follow this sequence:
- Add one or more projection folders.
- Leave the Glob pattern as tomo_*.tif unless your files use a different name.
- Save the output into a child folder.
- Use Mode = auto.
- Start with Correction = algotom.
- Keep the default parameters for the first run.
- Run the job.
- Read the log and compare the corrected results.

Default GUI values referenced in this manual: Correction = algotom; Workers = 12.